| Ball State University Environmental Health and Safety Office Laboratory (Self) Inspection Form |  |
| --- | --- |

It is the goal of the Environmental Health & Safety Office (EHS) to work cooperatively with Principal Investigators, faculty, and laboratory workers to achieve compliance with University safety policies, the Chemical Hygiene Plan, Biological Safety Plan, other BSU policies, and governmental regulations.  Unsatisfactory items and laboratory safety violations should be corrected or addressed through proper planning as soon as possible.  Issues that represent an immediate or imminent hazard to University Personnel, risk to the environment, or potential to cause damage to University facilities are classified at Category 1 Deficiencies (highlighted in red) and are to be corrected immediately.  Superscript numbers associated with inspection items correspond to the resources listed on the back page of this Form, the complete documents for which are available through the EHS website at http://cms.bsu.edu/about/administrativeoffices/riskmanagement/ehs/envhealth/labsafety  Contact the EHS Office with any questions.

| Department: |  |  |  |  |  |  |  |  | Date: |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Building and Room Number: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab Supervisor/PI Name: Email: Phone: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Person(s) Interviewed: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspector(s): Title / Dept.: General Laboratory Use Information. Chemical, Biological, or Physical Hazards Present in Laboratory, Clinical Area, or Stockroom. Extremely Hazardous Substances12; Chemicals of Interest 8; Select Agents 29, 31. Water Reactives; Pyrophorics; Explosives or High Energy Reactives. Carcinogens*25; Reproductive hazards*25; Acute toxins*25. Flammables; Corrosives; Formaldehyde. Peroxide formers 27; Benzene; Acrylonitrile. Methylene chloride; Cryogenic materials; Nanoparticles or UFPs. Laser(s) 23; X-ray machines 32; Radioisotopes. RG1 (BSL1 work) Organism(s) 28; RG2 (BSL2 work) Organisms(s) 29; Infectious Wastes/Materials(s) 1, 29. Research Performed: Yes No, if Yes, under the auspices of the: (* Particularly Hazardous Substances List 25). IRB; IACUC; RSC. IBC; Other-specify:; Protocols Available: Yes No. Room or Space Use. Research Laboratory; Stockroom; Classroom lab. Preparation Room; Storage; Other: ______________________. Laboratory animal care; Separate Administrative Space Available for Use: Yes No. |
| A. Administrative and Safety Basics |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Access to Safety Data Sheets (SDSs)30 and Pathogen Safety Data Sheets (PSDSs)26 in Lab? Personnel must have ready access to current SDSs for all hazardous chemicals in the laboratories and PSDSs for any infectious organisms. These may be managed as printed hard copies in an organized fashion such as a binder. SDSs /PSDSs may also be maintained through a bookmarked Internet site or as an electronic file. A functioning computer with internet access and a functioning printer must be then be available in the laboratory. All laboratory personnel must know where and how to access the safety data sheets (SDS) for the hazardous materials or infectious materials they use in their operations. Laboratories are strongly urged to print the SDS sheets for their chemicals from the manufacturer that produced them and keep them in a clearly marked three ring binder in the laboratory on a bookshelf or other organized fashion where they will be visible to all researchers. These printed SDS must be updated and current. Be aware that any MSDSs (Material Safety Data Sheets) present are, by definition, obsolete. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Hand Washing Sink, Hand Soap, Paper Towels, and Waste Container Available to Assure Proper Sanitation and Decontamination? On leaving the designated area or the laboratory area, remove protective apparel and thoroughly decontaminate, disinfect, or dispose of contaminated items. Thoroughly wash hands and forearms with soap and water. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| No Smoking, Eating, Drinking? 17 Smoking, Eating, Drinking, and the application of cosmetics are prohibited in all University laboratories. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Unauthorized Occupants? Laboratory workers must limit access by unauthorized occupants and challenge or report any suspicious activity or people to the University Police. Children should not be allowed in the laboratory. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab Secured? Assure that the laboratory is kept closed, and locked when no one is present in the laboratory. Maintain vigilance in protecting chemicals, biologicals, and equipment. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Emergency posting current, emergency contacts posted, Laboratory sign is accurate? 16 Is the laboratory hazard sign accurate? Is an up to date list of emergency contacts and phone numbers posted or listed near the laboratory hazard sign? Are the hazard classes present clearly identified? |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Chemical Inventory Available? 4 An up-to-date chemical inventory must be physically maintained in the laboratory (normally with the Safety Data Sheets) or through ChemTracker for all hazardous chemicals located in the laboratory. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Biological Material Inventory Available? 1, 29 A biological agent inventory must be maintained. All RG I organisms must be identified and an inventory submitted to the University Biosafety Committee (UBC), and all RG II or III organism use should registered with the UBC on the approved forms. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Chemical Hygiene Plan and Task Specific Training? 4 All members of the lab must have Chemical Hygiene/Safety training. Electronic management of training records is encouraged. Laboratory personnel must receive general, followed by lab-specific task and chemical specific training on all processes, chemicals, equipment, and hazards in the laboratory. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Biological Safety Plan and Lab-Specific Training? 1, 29 All members of the lab must be proficient in the BSU Biological Safety Manual procedures and precautions and, if RGII or III organisms in use, a risk-assessment must have been performed and a Lab-Specific Biological Safety Manual must be available. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Written Safety Procedures (SOPs)? 32, 4 Written Standard Operating Procedures are required for all procedures, protocols, lessons, and use of hazardous chemicals or biological agents. These written procedures should discuss the hazards involved in the operation and ways to mitigate the hazards (Risk Assessment). The written procedures should dictate the types of personal protective equipment, fume hood use and other safety practices. They should be revised whenever there is a change in procedure or a new procedure or chemical is added. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab Doors and Windows Closed? 15 Lab doors should remain closed at all times to assure optimum laboratory ventilation. Keeping lab doors closed also minimizes odors and helps prevent the spread of smoke in the event of a fire. Windows in doors must not be covered or obscured such that it prevents visibility into the laboratory from the hallway or adjacent rooms. Means to rapidly cover the windows should be available in event of an active shooter alert. Ventilation supply and exhaust ducts must not be blocked or impeded. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Emergency Training? 5 Laboratory members should know the procedures to follow for a laboratory accident; i.e. use an emergency phone or dial 5-1111 or 911 for help, location of first aid kits, eye washes, accident reporting procedures. Researchers need to develop procedures to make operations safe should a ventilation or power failure occur. Lab workers should know where the safe emergency evacuation gathering points are located. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| B. Electrical 22 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| General Condition? Check for frayed power cords, unsafe electrical operations, open/exposed wires, grounding plugs removed, static discharges, etc. Adequate electrical service is necessary in the lab to avoid the use of unsafe practices such as permanent use of extension cords. Electrical equipment must be maintained in good condition. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Use of Extension Cords? Extension cord use should be eliminated. Contact BSU Work Control for installation of grounded receptacles. Check to be sure that no cords are frayed. Multi-plug devices are allowed provided they are UL listed with a built-in circuit breaker and used in accordance with the manufacturer’s intended use. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Breaker/Circuit Identification and Access? Circuit breakers that service laboratory equipment should be identified as such. Electric service panel access must be maintained. A 30-inch wide and 36-inch deep floor space and area in front of the cabinet must be kept clear to allow access and safety from arc flashes. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Ground Fault Protection? Outlets located near sinks (within three feet) or other sources of water should be on a ground fault circuit or otherwise ground-fault protected. Cheater plugs not used and equipment grounded as necessary. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C. General Safety 17, 22 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Housekeeping, Surplus Equipment and Egress? Housekeeping must be maintained so that the aisles are clear to allow for emergency egress. A minimum of 36” of clear aisle space in required throughout the lab. Could a person exit the lab quickly without tripping over objects? Materials, chemicals and debris should not be stored on the floor or in the path of egress. Excess or surplus equipment and supplies should be disposed or relocated to a storage location. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Hot Surfaces/Equipment? Hot surfaces or equipment should have integral warning lights, be posted with an appropriate warning sign, or have some mechanism to indicate that the surface is hot. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Safety Shower/Eye Wash? Safety showers and eyewash units should be located in the lab or nearby (within 10 seconds) and accessible at all times. Eyewash must be in the laboratory if corrosive, irritants, or particulate materials are used. They should bear a tag indicating an inspection within the last year by the EHS Office. Eyewash units should be tested and flushed weekly for 3 minutes by lab occupants, faculty, or the PI. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Refrigerators/Freezers? Laboratory refrigerators and freezers must not be used to store of food or drink for human consumption. Refrigerators and freezers used to store flammable, combustible and corrosive chemicals must be approved and certified. Units must be signed or labeled as NOT FOR FOOD USE. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Needles/Syringes Secured? Needles and/or syringes must be kept secured at all times. This means kept in a locked drawer or cabinet or in a locked laboratory and under surveillance of lab personnel when the lab is unlocked. Do not leave uncapped needles in fume hoods or on laboratory benches. Rigid properly labeled biohazard containers must be available and used for sharps disposal and replaced when ¾ full. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Tripping/Slipping Hazards? 22 Floors should be in good repair, i.e. no tripping hazards caused by mats, cracks, holes, protrusions, missing tiles, ethernet cords, wires, hoses, tubing, etc. Chemicals must not be stored unprotected on floors. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Potable Water Protection? Hoses connected to sink faucets should not extend below the plane of the sink surface or proper back flow prevention devices should be installed. Lab sinks should be equipped with backflow prevention. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| First Aid Kits? There should be a first aid kit available in case of minor injuries. It should remain stocked, readily located, and be accessible during the hours of operation of the lab. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Ice Machines/Microwave Ovens (Labeled)? Laboratory microwaves or ice machines must not be used with food or drink for human consumption. Make sure they are signed appropriately. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| D. Fire Safety Concerns |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Detectors/Sprinklers? If there are smoke/heat detectors, audio/visual devices or sprinklers in the lab, make sure that nothing is stored near them that would interfere with their intended operation. No storage within 18 inches of sprinkler heads. Where no sprinklers, no combustible storage closer than 2 feet from ceiling. Visually check for physical damage or leaks. Report any deficiencies to Work Control. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Combustible Storage, such as excess boxes, shipping containers, paperwork? Check for excess storage of combustibles or combustibles near any hot surfaces or equipment. Maintain at least 18 inches from hot surfaces or equipment. Eliminate unnecessary storage in the lab space. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Special Concerns? If drying operations are performed in the lab, are there procedures written for safe operation; i.e. are lab personnel instructed not to use combustible trays for holding materials to be dried? Limit the use of Bunsen burners. Develop written safety procedures for all heating operations and train lab workers. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Fire Extinguishers? Fire extinguishers should be located in each laboratory for use by trained personnel only. They are inspected monthly by EHS. Check to make sure the extinguisher is present, charged, and is not damaged. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| E. Chemical Concerns 13 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Storage by Hazard and Chemical Storage? 6 Chemicals need to be stored by hazard class, i.e. Reactives, Flammables, Poisons, Oxidizers, Inorganic Acids, Organic Acids, Caustic Materials, etc. Organic acid should not be stored with inorganic acids. Carcinogenic chemicals, acutely toxic chemicals and reproductive toxics should be used only in established Designated Areas. Pay close attention to special storage requirements such as refrigeration needs, desiccated, inerted, etc. Liquids must not be stored above shoulder level. Open shelving units should preferably have a lip. Hazardous chemicals must not be stored on the floor. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Flammables/Combustibles? Flammables and combustible materials should be stored in flammable storage cabinets wherever possible and always kept away from ignition sources. Lab safe or explosion-proof refrigerators and freezers are required when flammables and combustible chemicals must be stored at refrigerated temperatures. Regular/residential refrigerators and freezers should bear the caution statement prohibiting storage of these materials. No more than 10 gallons of flammables outside of approved safety cabinets. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Explosives and Pyrophorics? Materials classified as “explosives” under DOT or IFC are not allowed in university laboratories. Pyrophorics are allowed only in building equipped with automatic fire-suppression systems. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Peroxide formers27, such as ethyl ether and tetrahydrofuran, dated when the material was received and when opened. Material is stored past the manufacturer's expiration date. Date all peroxide-formers upon receipt and upon opening. Unless an inhibitor has been added by the manufacturer, materials should be properly disposed of after 18 months from the date of receipt or 3 months from the date of opening. If an inhibitor is added, dispose of the material based on the manufacturer's expiration date. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| High Toxins and Corrosives? 13 Acute toxins and corrosives, as well as flammables, explosives, etc., should be stored within specific fire control areas only in quantities beneath the threshold requiring a High Hazard Occupancy rating under the International Fire Code. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Excessive Quantities? Quantities of chemicals kept in the lab should not be excessive. Outdated or expired chemicals should be disposed of using proper disposal methods outlined in the BSU Waste Management Guide. Legacy chemicals or any beyond 2 years in age should not normally be found. Observation or storage of chemical containers on floors, in fume hoods, working counters--unlabeled, aged, or deteriorated containers, duplicate open containers, etc., is indicative of excess chemicals in the laboratory. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Compatible Containers and Kept Closed? 7 All containers holding hazardous chemicals or infectious materials must be kept closed at all times other than when filling or emptying. The container materials must be compatible with the chemicals contained and in good condition without evidence of leakage. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Chemical Transfers? 7 Transfers of flammable liquids between containers should be grounded and bonded as necessary. Mouth pipetting is prohibited. Hazardous chemical transfers should not be done at sinks or near drains and are preferably performed within or above secondary containment. Ensure adequate ullage available in receiving container(s). |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Container Labels? 20 All primary (manufacturer) chemical containers must have the original label legibly affixed to the container. Primary (manufacturer) labels must be replaced if lost, damaged, or illegible. Secondary (working or transfer) containers must bear at least the common name of the chemical and appropriate hazard labels. Reaction flasks, samples and other lab containers must be labeled as well. Common chemical names or IUPAC nomenclature should be used whenever possible. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Chemical Spill Kits? 5 Chemical spill kits must be readily available and, where necessary, contain absorbents or neutralizing agents for certain specific hazardous materials (e.g., aldehydes, hydrofluoric acid) if present. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Particularly Hazardous Substances (PHSs) 25: Reproductive Toxins, Highly Toxic Chemicals, Or Carcinogens require specific SOPs, training, and designation of areas of use. See the BSU Chemical Hygiene Plan for identification of these substances. All work with PHSs require particular attention and adherence to strict standards under the OSHA Laboratory Standard – Tables in the CHP identify these substances. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab limits the use of mercury thermometers and metallic mercury? Laboratories should eliminate the use of metallic mercury. Replace all mercury thermometers with digital or alcohol thermometers. Eliminate metallic mercury manometers, psychrometers, and sphygmomanometers (blood pressure units). Contact EHS for proper disposal and recycling of the units and equipment. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| F. Waste Management 18, 19, 11 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Quantity of Chemical Waste Accumulated? Chemical waste should be identified with LAB WASTE labels. Do not accumulate wastes for longer than 6 months for or accumulate over 55 gallons of chemical waste materials in a lab before removal. Full waste containers must be removed promptly. See the BSU Laboratory Waste Management Plan. Contact EHS at 5-2807 if you need assistance or removal of wastes. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Properly Labeled? Abbreviations or trade names must not be used to identify contents on chemical LAB WASTE labels. Common chemical or IUPAC nomenclature must be used. Label every constituent added to the container with approximate volumes in percent. Unknown substances, explosives, pyrophorics, and some highly reactive chemicals are forbidden in transport and will be treated and disposed at the department's expense. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Properly Stored/Secured? Liquid waste must be stored in a secondary container, segregated by chemical compatibility. Do not fill liquid waste containers over 90% full. Waste containers (both liquid and solid) must be capped or sealed at all times unless material is actually being added. Infectious waste boxes must be constructed properly with packing tape and lined with red Biowaste bags, or rigid containers with autoclave liners utilized. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Proper Waste Containers Available? 7, 18, 19 Waste containers must be compatible with contents and kept closed and properly identified. Biological waste, broken glass, and sharps must be stored in approved containers. Contact the EHS Office if approved containers are needed. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Evidence of Improper Waste Disposal? Normal trash cans or recycle bins should not contain inappropriate materials; i.e. grossly contaminated PPE, materials that can be construed as chemical or biological waste. Sinks should be free of stains and waste containers should not be stored in close proximity to sinks unless they are in secondary containers. Broken laboratory glass is to be collected in "glass only" boxes. Chemical bottles must be triple rinsed and the labels defaced or removed prior to disposal or recycling. All biological waste must be disposed of properly in infectious waste boxes or sharps containers, or autoclaved as necessary prior to disposal. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab personnel have received chemical waste management training? 19 At least one person in each laboratory (or the department laboratory manager, or equipment manager, or lab technician) where chemical wastes may be generated must attend hazardous waste training annually. Waste management must be included in the laboratory general or lab-specific training programs and be documented. Contact EHS for more information. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Biological wastes 1, 19 Potentially infectious wastes, including Bloodborne pathogen-regulated wastes, must be properly contained in tight containers with inner impermeable bags, kept closed or covered, and labeled with contents. All such wastes must be autoclaved prior to disposal as solid waste in plain outer bags or be disposed in infectious waste containers or sharps containers for proper off-site disposal. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| G. Physical and Environmental Concerns 22 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Moving Parts Guarded? All belts, blades or other moving parts and nip-points on equipment should be guarded or otherwise protected. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Shields Used? Adequate blast shields should be used when conducting experiments that could react violently or explode, or if there is the possibility of a flying projectile or object. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment Clean and Operable? All lab equipment should be clean and in good working order. Damaged/out of service equipment must be tagged and removed from the lab for proper servicing, repair, or disposal. Equipment and devices not in use should be excessed or stored elsewhere than in an active laboratory |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment and Analytical Process Instructions and Training? Operation manuals must be available for all analytical and preparatory equipment and operator training on any equipment staff/students are to use should be documented. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Sharp Edges, Points? Check for sharp edges or points sticking out on equipment, furniture etc. that could cause an injury. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Laboratory Environmental Conditions – Wet Floors, Temperature Concerns, Wet/Stained Ceiling Tiles, Mold, Peeling Paint Call Facilities Work Control or the EHS Office to report issues with the temperature, wet/stained ceiling tiles and mildew. Use matting in areas where the floors periodically get wet such as in front of -80 Freezers. Lab staff must be proactive in reporting problems that present hazards and need to be corrected. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Oil mist filters installed on oil filled vacuum pumps? Oil mist filters maintained properly? No oil on the floor, in the exhaust lines or no other indications of a leak or issue. Used oil kept free of solvent contamination and labeled for recycling or handled as LAB WASTE if contaminated by solvents. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab Space and Furnishings Adequate for Activities? Room for safe research or educational activities: Benches, aisle space, lighting, storage capacity, separate administrative area, local and general ventilation adequate. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Utility Service and Protection? Gas, water, electric, and waste drain piping and service in good repair, valves work, no corrosion, emergency shut-offs available, piping contents identified, and accessible. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| General Environmental Conditions? Humidity is 20-60%, temperature is 68-76̊ F, negative air pressure relative to hallways, acceptable CO and CO2 levels, lighting adequate. HVAC vent diffusers clean, vents not blocked. Problems with allergies or respiratory distress, itching eyes, etc. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| H. Personal Protective Equipment 21, 1, 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Proper Lab Attire, Personal Protective Equipment (PPE) Available? 21 In general, proper lab attire will consist of safety glasses, lab coats, gloves, long pants that cover the legs and closed-toe shoes. Face shields, splash goggles, hot mitts, cryogenic gloves, and specialized PPE should be provided and available for particular hazards. Fire resistant lab coats required if working with highly flammable or pyrophoric materials. Respirator users participate in BSU Respiratory Protection Program. Laboratory-specific PPE requirements posted and restrictions/exclusions policy enforced for non-compliance. Departments to provide required PPE to staff and to assure that students provide required PPE for their activities or exposures. The specific PPE to be employed is to be based on a hazard assessment. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Proper PPE Being Utilized 21 All personnel and students in active laboratory work setting wearing the minimum designated PPE for the laboratory and additional or specific PPE for the particular tasks being performed. PPE training records available for staff and students. Efforts must be made to minimize all exposed skin. Sandals, short skirts or shorts and sleeveless shirts are not appropriate for the lab. PPE should not be worn in non-lab areas. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Safety Glasses Being Worn? Eye protection (of varying levels) is mandatory in all University laboratories except laboratories used exclusively for computers. (See Policy in the BSU CHP) |  |  |  |  |  |  |  |  |  |  |  |  |  |
| I. Ventilation and Engineering Controls 2, 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Fume Hood Operation – Check Before Use! 3 The lab manager, instructor, or researcher should make sure the fume hood(s) are operating properly before use. Check the automatic alarms or flow indicators. Make sure the sash is at proper height and, the baffles are properly set based on the vapor density of the chemicals to be used. Verify the fume hood is operating properly by holding a ribbon or tissue (Kim Wipe) below the sash and observing flow. Ensure lab doors are closed and personnel traffic past the hood is minimized. Notify EHS or Work Control of any concerns. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Proper Use of Fume Hood? 3 Fume hood work should be performed approximately six inches into the hood, i.e. not right at the front edge. The sashes should be positioned to produce the flow rate indicated on the certification sticker. They should not be used for storage of chemicals. Large equipment must be raised about 1 inch from the fume hood deck. Only chemicals, materials and supplies for the current experiment or operation should be in the fume hood. Fume hoods must be kept clean and should be decontaminated on a periodic basis. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Fume Hoods Available? 4 If operations in the lab involve materials which present an inhalation exposure hazard, a fume hood, exhaust trunks or biosafety cabinet must be available for use. These operations include, but are not limited to, work with carcinogens, reproductive toxins, acutely toxic materials, flammable liquids, materials with high vapor pressures, dusts, mists, volatile materials, soldering operations, welding operations, or with any chemical that has an occupational exposure value < 50 ppm or emits nuisance odors. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Fume Hood Inspected/Flow Measured?4 Fume hoods and local exhaust trunks are inspected by the EHS for proper operation twice per year. An inspection label should be present to verify this inspection. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Biosafety Cabinets? 2 All Biosafety Cabinets must be inspected and certified by a contractor annually. Chemicals should not be used in cabinets. Bunsen burners should not be used in biosafety cabinets. Biosafety cabinets should be positioned away from doorways and supply/exhaust vents; this will disrupt the laminar flow. Procedures for disinfection of the cabinets must be available. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| J. Compressed Gases 9 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Proper Storage? Gas cylinders should not be stored in laboratories unless they are being used for the current operation. Empty cylinders should be moved out as soon as possible. Do not store a gas cylinder between a user and an exit. Compressed gases with a health hazard rating of 3 or 4 must be stored in a rated ventilated gas cabinet. Compressed gases with a fire hazard of 4 (hydrogen, for example) require a second means of egress form a laboratory or must be stored in a rated ventilated gas cabinet. Contact EHS for additional assistance with compressed gas safety. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Proper Labeling? Gas cylinders must be labeled with the contents. Unknown gas cylinders are very hazardous and expensive to dispose. Empty gas cylinders must be so labeled and remain secured. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Cylinders Capped or Restrained properly? Caps should be kept on the cylinders when not in use. Cylinders need to be restrained by chains, belts, or bars at approximately two-thirds height from the floor. The cylinder must be of good quality and firmly attached to a structural member. Bench clamps should be avoided and only used if a gas cylinder bracket cannot be permanently mounted. No more than 2 cylinders may be secured together. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Proper Lines and Regulators? Gas supply lines need to be compatible with the gas being used. All gas cylinders must have an appropriate regulator. Do not use thread tape on the regulator/cylinder connection(s). The suppliers can provide information about proper line material. Regulators should be replaced or recertified on a regular basis. Contact the gas supplier with questions. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Facility/Gas Safety? Storage of cryogenic, liquefied, and other significant compressed gas storage should be evaluated relative to room size, and gas storage/expansion capacity for the need for oxygen, combustibility monitors and other features such as purge fans or blast panels. Proper precautions and protective devices must be provided. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab personnel trained on compressed gases and gas cylinder safety? Laboratory personnel must receive task and chemical specific training on all processes, chemicals, equipment and hazards in the laboratory and Standard Operating Procedures should be available. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| K. Security 15, 16 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Lab Access Doors and Passages Secured? Doors are locked when no persons are present in the lab, prep room, or storage rooms to prevent unauthorized entry. Key cores, latches, and closing mechanisms are functional and doors operate properly. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Key and Access Controls? Written departmental and lab-specific key control procedure is effective to limit access to only authorized personnel. Current records of access and key possession are available for inspection. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Intruder Precautions? Individual room and doors and passages and building access can be secured from the interior. Door windows can be rapidly and temporarily covered to prevent viewing of interior. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Emergency Procedures?22 Fire doors remain closed except when entering/exiting. Laboratory personnel and students are trained in emergency procedures (fire, explosion, chemical/biological spills or exposures, active shooter, natural disasters) recognition and response (evacuation routes and rally points). Drills or safety meetings are held to verify knowledge of emergency procedures. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Access Limitations? Written policy prohibiting working alone with hazardous materials, infective biological agents, or hazardous operations, exposure to reproductive toxins or hormone disruptors, and access limitations to the lab spaces after normal hours exist, are communicated, and enforced. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| L. Administrative Controls/Work Practices Observations 17, 22 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Administrative Tasks Not Performed in Laboratory Spaces Separate physical (and ventilation) space available for staff/student use outside of the active laboratory area, equipment, or benches, where exposure to chemical or biological agents would be a concern.. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Safety in University Laboratories Brochure17 All staff and students familiar with the “Safety in University Laboratories” brochure provided by the EHS Office that summarizes safe laboratory procedures and prohibitions in laboratories, studios, and workshops. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Demeanor 17, 21 No disruptive behavior presented or allowed in laboratory. Staff and student attire, accoutrements, and jewelry are appropriate to the laboratory activities, engineering controls in use, and required personal protective equipment (PPE) relative to the hazards presented. This may require removal of jewelry, restraining loose hair, and suitable clothing that minimizes skin exposure. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Training All personnel and students have received general laboratory safety training and lab-specific training for the particular laboratory, chemical and biological agents, equipment and processes. Familiar with the BSU Chemical Hygiene Plan and Biological Safety Plan, and PPE Programs. Lab-specific safety training provided. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Hygienic Practices Hand washing, cleaning and disinfection of work surfaces and equipment performed as necessary. Visible contamination of spillage not observed. No evidence of eating, drinking, chewing gum, smoking, applying cosmetics, etc. in the lab spaces. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Inspection Item |  |  |  |  |  |  |  | S |  | UNS | N/A |  | Comments |
| Awareness Staff and students are aware of the location and meaning of container labels, hazard warnings, SDSs, and the location and use of all emergency equipment – eyewash/showers, fire extinguishers, exits, utility shut-offs, spill response and first-aid supplies, emergency exit routes. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Local Exhaust Systems Chemical fume hoods, biosafety cabinets, and local exhaust systems used when necessary, checked for function, used properly (sash heights, proximity to intake, etc.). SDS restrictions on chemical use in ventilated areas followed. Any filters changed as needed or on schedule. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Chemical Hygiene Plan, Biological Safety Manual, and SDSs, or PSDSs 1, 4 Chemical Hygiene Plan and Biological Safety Manual available for reference (Lab-specific Manual if required). Safety Data Sheets (SDSs) readily available for all hazardous chemicals in the lab space. Pathological Safety Data Sheets (PSDSs) available for any biological agents. PSDS may be obtained through the Public Health Agency of Canada at https://www.canada.ca/en/public-health/services/laboratory-biosafety-biosecurity/pathogen-safety-data-sheets-risk-assessment.html |  |  |  |  |  |  |  |  |  |  |  |  |  |
| SOP Development and Availability 32 Standard Operating Procedures (SOPs) for the lessons, demonstrations, procedures or research being performed are available, including hazard assessment and controls and step-by-step instructions for safety and effectiveness of the procedure. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Equipment Use Energized, water, compressed air, or gas equipment not left unattended unless safety measures covered in SOP or protocol, equipment instructions followed, cleaned, and shut down properly after use. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Personal Articles Provisions for storage of student backpacks, coats, cell phones, and other personal items in a manner or location to prevent contamination and the possibility of theft. Sealing of exposed articles (lab coats or gowns, electronic devices) in plastic baggies may be necessary or advisable. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Occupational Health and Safety28 Staff aware of occupational health and safety provisions for employees working in laboratories – medical evaluations and consultation availability, exposure assessments, etc. Aware of emergency, exposure, accident, and medical emergency procedures. Advisory policy on pregnancy and reproductive toxins made available. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Security Unauthorized personnel or student access prohibited. Laboratory doors locked when no one present in the room(s). Chemical and biological storage and inventory monitored. Working alone in laboratory not allowed or only subject to specific conditions. Any highly reactive materials, toxins, etc., precursors, etc., kept in locked cabinet or secured area. Equipment secured to prevent theft. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Culture of Safety Evident Faculty, PI, and Staff recognize laboratory safety as an integral component of the university science and arts practice and educational goals. Exhibit and promote proper safety practices, PPE use, and behavior in laboratory, setting a good example for other staff and students. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Reporting of Physical Problems Faculty, PI, and Staff report facility physical defects (leakage, corroded piping, electrical deficiencies, defective alarms, peeling paint, chemical and biological exhaust systems failures, etc., in the laboratory, studio, or stockroom promptly to department administration or university Work Control. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Working Alone The laboratory and management have policies limiting working alone with hazardous chemicals, processes, or infectious organisms in the lab. Staff and students are aware of university policy advisory regarding pregnant persons and the potential hazards of exposure to hazardous chemical, radiation, and biological hazards in the lab. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| M. SPECIFIC REGULATORY PROGRAMS – Note Applicability or Refer To Appropriate Office. Inspection Item; S; UNS; N/A; Comments. Radioactive Materials (NRC License) – refer to BSU Radiation Safety Officer (5-8066). 1; All areas/containers with radioactive materials are identified with signs/labels.. 2; Material is secured when lab is unattended.. 3; Material is identified and shielded in the laboratory. Lasers –refer to Laser Safety Officer 23. 1; Class IIIB or IV lasers are registered and signage is adequate.. 2; SOPs for Class 3b and 4 laser use. 3; Operators trained in SOP and ANSI Laser Safety Standard. X-Ray Equipment – refer to X-Ray Equipment Radiation Officer (5-2807) 32. 1; XRFs, TEMS, SEMS, etc. are registered,. 2; Dosimetry and records provided where needed. 3; Manufacturer or in-house training documented for all operators. 4; Required and appropriate signage provided at access doors. Controlled Substances – refer to EHS Coordinator (5-2807) 10, 24. 1; Researcher/Teaching Registration and Licenses (ISBP and DEA). 2; Scheduled Material is secured – access limited to authorized users. 3; Required records and inventory maintained. 4; Listed DEA substances cleared through EHS Office and documented. Hazardous Materials Shipping and Receiving – refer to EHS Coordinator (5-2807) 14. 1; Persons shipping or receiving Hazardous Materials DOT trained. 2; Training and shipping records (manifests, shipping papers) available. Plant Pest and Noxious Weed Control Program 30. 1; Compliance with USDS Animal & Plant Health Inspection Service Permit, shipping, and record-keeping. Animals (laboratory, demonstration, or exhibition) 22. 1; Compliance with USDS/PHS/OLAW/IACUC animal care requirements – facilities, training, housing, OHSP, record-keeping. Biological Safety Facilities or Research – refer to EHS Coordinator (5-2807) 1. 1; Level 1 2 3 Risk Group Organisms in use or present. 2; Level 1 RG organisms Inventory and Level 2/3 Registered with UBC?. 3; Biological Safety Manual Compliance. |
| N. Additional Inspection Comments or Concerns (list and describe) |  |  |  |  |  |  |  |  |  |  |  |  |  |
| O. Corrective Actions |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Items marked as unsatisfactory will be reported to the Department Chair, and to the responsible PI or Faculty. Category 1 Deficiencies must be rectified as soon as possible. Category 2 Deficiencies should be rectified in 30 days or sent to the Department Chair for planning action or assistance. Deficiencies were communicated to: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Name: |  |  |  | Signature: __________________________ |  |  |  |  |  |  |  | Date: |  |
| Name: |  |  |  | Signature: __________________________ |  |  |  |  |  |  |  | Date: |  |
| Date(s) Items will be Rectified: Corrective Action Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| P. Follow-Up Inspection |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A follow-up inspection to determine if agreed upon steps to rectify outstanding issues were implemented was completed: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| By: |  |  |  |  |  |  | Date: |  |  |  |  |  |  |
| Unsatisfactory Items Remaining (Listed on right) |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Follow-Up Inspection Notes |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Name: |  | Title: |  |  |  | Date All Confirmed: |  |  |  |  |  |  |  |

References and Resources (Alphabetical: Keyed by Number to Inspection Form)

| No. | Reference Document |  | Description |  |  | Source |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Biological Safety Manual |  | Establishes Biosafety Committee and procedures for research involving infectious agents, rDNA, and synthetic NA to conform with CDC/NIH standards and BSU EHS guidelines. |  |  | NIH EHS |
| 2 | Biosafety Cabinets |  | Brief on safety standards for Biological Safety Cabinet use |  |  | OSHA |
| 3 | Chemical Fume Hood Safety |  | Brief on safety standards for Chemical Fume Hood use |  |  | OSHA |
| 4 | Chemical Hygiene Plan |  | OSHA-required overall plan for chemical safety in all BSU academic laboratories including policies, procedures, prohibitions, tables, etc. |  |  | EHS OSHA |
| 5 | Chemical Spill Control Plan |  | Response plan covering preparedness, prevention, and control of chemical and hazardous material releases |  |  | EHS EPA |
| 6 | Chemical Storage Segregation |  | System for segregating incompatible chemicals in storage |  |  | Stanford U. |
| 7 | Chemical Waste Containers |  | Discusses proper containers and storage of chemical wastes |  |  | EHS |
| 8 | Chemicals of Interest |  | Homeland Security List of chemicals with quantities for screening of potential chemical weapons. |  |  | DHS |
| 9 | Compressed Gas Cylinder Safety |  | Fact sheet on safety precautions in storage and use of gases |  |  | OSHA EHS |
| 10 | Controlled Substances List |  | Alphabetical list of DEA Controlled Substance Schedules |  |  | DEA |
| 11 | Drain Disposal of Lab Wastes |  | Discusses acceptability/prohibitions in drain disposal of wastes |  |  | EHS |
| 12 | Extremely Hazardous Substances |  | List of chemicals and threshold quantities that then require community reporting under Community Right to Know regulations. |  |  | EPA |
| 13 | HazMat Allowable “Exempt” Storage Tables |  | International Fire Code “exempt” quantities of hazardous materials allowed in any “Fire Control Area”-further limited by floor level. |  |  | IFC |
| 14 | HazMat Shipping Flyer |  | Advises all employees of DOT HazMat shipping requirements |  |  | U.S. DOT |
| 15 | Lab Door Hazards |  | Bulletin regarding need to keep lab doors closed and effective |  |  | OSHA IFC |
| 16 | Lab Door Sign Instructions |  | Instructions for preparing lab door hazard signage |  |  | EHS |
| 17 | Lab Safety Handout |  | BSU pamphlet covering basic required laboratory safety procedures |  |  | EHS |
| 18 | Lab Waste Management Plan |  | BSU’s EPA-required laboratory chemical waste management program |  |  | EHS EPA |
| 19 | Lab Waste Management Training |  | PDF file of Laboratory Waste Management Training PowerPoint presentation |  |  | EHS EPA |
| 20 | Labeling of Chemical Containers |  | Fact sheet on OSHA requirements for identity labeling and hazard identification for chemical containers. |  |  | OSHA |
| 21 | Laboratory PPE Guide |  | Guide for PPE considerations or requirements for laboratories |  |  | EHS |
| 22 | Laboratory Safety Guidance |  | OSHA guide to their laboratory regulations, standards, and related laboratory safety requirements and practices. |  |  | OSHA |
| 23 | Laser Safety Manual |  | Establishes laser classification system, registration program for Class 3B and 4 laser systems and signage, training, and safety standards. |  |  | EHS |
| 24 | Listed Regulated Chemicals |  | List I and II chemicals with legitimate uses, but designated as possible precursors to Controlled Substances by the DEA. |  |  | DEA |
| 25 | Particularly Hazardous Substances (PHSs) |  | Carcinogens, reproductive hazards, and toxins OSHA calls Particularly Hazardous Substances, requiring SOP development, Designated Areas. |  |  | OSHA |
| 26 | Pathogen Safety Data Sheet |  | Example Pathogen Safety Data Sheet – Candida albicans |  |  | CPHS/NIH |
| 27 | Peroxide-Former Handout |  | Precautions on handling and identifying peroxide-formers |  |  | EHS |
| 28 | Pregnancy and Lab Exposures |  | Fact sheet with precautions for pregnant workers in laboratories |  |  | EHS |
| 29 | Risk Groups and Select Agents |  | Infectious agent Risk Groups, defined by NIH, relative to Biosafety Levels. Also Select Agent, and Plant Pathogen listings & discussion |  |  | CDC/NIH DHS/USDA |
| 30 | Safety Data Sheet Example |  | Acetone Safety Data Sheet (example of SDS content) |  |  | OSHA |
| 31 | Select Agents and Toxins |  | Biological agents and toxins with the potential to pose a severe threat to public health and safety, animal and plant health, or products. |  |  | DHS USDA |
| 32 | SOP Form (Fillable in CHP |  | Fillable pdf (in CHP) for creating Standard Operating Procedures |  |  | EHS OSHA |
| 33 | X-Ray Safety Plan |  | Covers analytical X-Ray Machine registration, dosimetry, training, and safety procedures – electron microscopy, XRFs, etc. |  |  | ISDH |
| CDC |  | Centers for Disease Control |  | EPA | U.S. Environmental Protection Agency |  |
| CPHS |  | Canadian Public Health Agency |  | IFC | International Fire Code (Indiana adopted by reference) |  |
| DEA |  | U.S. Drug Enforcement Administration |  | ISDH | Indiana State Department of Health |  |
| DHS |  | U.S. Department of Homeland Security |  | NIH | National Institutes of Health |  |
| DOT |  | U.S. Department of Transportation |  | OSHA | U. S. Occupational Safety and Health Administration |  |
| EHS |  | BSU Environmental Health and Safety Office |  | USDA | U.S. Department of Agriculture (APHIS) |  |